## Sheet1
Row Labels | Sum of Count
Ask Ole | 40
PC-013-006 | 5
Data from unit was discarded | 5
PC-013-031 | 1
Data lost from device | 1
PC-014-113 | 2
No active analyser | 2
PC-021-009 | 5
Job type error | 5
PC-032-101 | 27
Read Error Correction dialog synchronisation error | 27
Cadance 100 issue - solved | 6
FM7DC-MFH-008 | 6
Internal error waiting for pipette. | 6
Check our time-out - should we extend time-out | 26
PC-049-000 | 26
Rinse job timed out | 26
Conveyor 4000 | 53
CV-CPU-018 | 10
Rack en face des capteurs. Faire un retour du rack | 10
CV-CPU-022 | 6
Mouvement de base rack | 6
CV-CPU-050 | 37
La taille du rack ne correspond pas | 37
Conveyor 7 Calibration | 1077
CV7-CV7-003 | 71
Rack size mismatch. Check if settings and sensor calibrations are correct | 68
Rack size mismatch. Check if settings are correct | 3
CV7-CV7-004 | 137
Rack movement error. Check if the movement is blocked | 137
CV7-CV7-009 | 869
Pipette or Stirrer counter error. | 869
Conveyor 7 Disconnect | 128
PC-014-127 | 128
Connection to Conveyor 7 is lost. | 128
Conveyor 7 Installation | 4
CV7-CV7-001 | 4
Setup data are invalid. Check if settings are correct | 4
Conveyor 7 Operator | 32
CV7-CV7-002 | 32
Rack detected in front of one or more sensors. Please rewind rack | 32
Conveyor 7 Other issues | 16
PC-021-013 | 16
Conveyor not rewound before restart | 16
Flow system Error | 8
PC-120-109 | 1
Baseline level is saturated. No or too little solution is passing through the flow cytometer. | 1
PC-120-110 | 7
Flow system error. Please begin by checking incubation valves. Other possible reasons could be valve VPL_V9, also check pipette for blockage. | 7
FM 7 DC Instrument | 1
FM7DC-RSH-011 | 1
Rinse-Sheath liquid temperature low. | 1
FM 7 Instrument | 1365
FM7DC-FCU-001 | 423
Needle inlet temperature too low | 423
FM7DC-IC-005 | 7
Intake bubble sensor (VPL_S1) calibration error. | 7
FM7DC-MFH-002 | 29
SSY motor timeout, sensor not found. | 29
FM7DC-MFH-007 | 36
Internal error waiting for sample intake. | 36
FM7DC-MFH-022 | 86
Sample intake failure (VPL_S1). | 86
FM7DC-MWH-006 | 6
Internal waste tank full. | 6
FM7DC-RSH-011 | 13
Rinse-Sheath liquid temperature low. | 13
FM7DC-RSH-014 | 6
Timeout when filling Rinse / Sheath unit. | 6
FM7-FCU-003 | 12
Laser detector receives too low signal from laser | 12
FM7-IC-001 | 1
Incubation module not found at startup. | 1
FM7-IC-005 | 229
Intake bubble sensor (VPL_S1) calibration error. | 229
FM7-MFH-007 | 2
Internal error waiting for sample intake. | 2
FM7-MFH-022 | 65
Sample intake failure (VPL_S1). | 65
FM7-MFH-025 | 1
Intake heater temperature low. | 1
FM7-MWH-003 | 1
Mix syringe motor error. | 1
FM7-MWH-006 | 179
Internal waste tank full. | 179
FM7-MWH-007 | 138
Internal waste tank sensor missmatch. | 138
FM7-MWH-008 | 4
Mix liquid sensor calibration error. Check supply. | 4
FM7-RSH-005 | 1
inse-Sheath pressure high | 1
FM7-RSH-006 | 91
Rinse-Sheath pressure low | 91
FM7-RSH-011 | 10
Rinse-Sheath liquid temperature low. | 10
FM7-RSH-014 | 22
Timeout when filling Rinse / Sheath unit. | 22
FM7-RSH-018 | 1
Rinse-Sheath valve 2 driver error. | 1
PC-120-112 | 2
Malfunctioning pinch valves | 2
Fossomatic 7 Baseline | 125
PC-120-111 | 125
Baseline out of limit. | 125
Fossomatic 7 Disconnect | 361
FM7DC-IC-010 | 20
Instrument unexpectedly reset. Sample under pipette might be contaminated. | 20
FM7-IC-010 | 159
Instrument unexpectedly reset. Sample under pipette might be contaminated. | 159
FM7-MFH-020 | 1
Internal communication error. | 1
PC-006-141 | 117
Instrument communication error | 117
PC-014-117 | 4
Device failed during device mode change | 4
PC-014-118 | 4
Device failed during device mode change (Standby) | 4
PC-014-119 | 26
Device failed during device mode change | 22
Device failed during device mode change (Auto) | 4
PC-014-124 | 24
Device unexpectedly been reset | 24
PC-014-127 | 6
Connection to Fossomatic 7 is lost. | 6
Installation | 10
PC-006-124 | 2
Cannot Initialize Gateways | 2
PC-006-140 | 2
No hardware driver found | 2
PC-014-115 | 1
Unable to connect to device(s). | 1
PC-071-050 | 4
Failed to verify that COM port is configured | 4
PC-120-108 | 1
Instrument Configuration must be carried out. | 1
LIMS Server Error | 160
PC-016-000 | 4
Export of result transmission has timed out. | 4
PC-059-101 | 154
External data communication server has closed the network connection | 154
PC-071-001 | 1
Communication failure on the CS83 on RS232 (Serial) data connection | 1
PC-074-201 | 1
An error occurred during database export | 1
MilkoScan 7 | 693
MSC7RM-CPU-001 | 111
Accumulator fill timeout | 111
MSC7RM-CPU-002 | 54
Accumulator leak | 54
MSC7RM-CPU-003 | 164
Accumulator normal timeout | 164
MSC7RM-CPU-008 | 1
H-pump stroke-timeout | 1
MSC7RM-CPU-009 | 81
Measure time too long | 79
Messdauer zu lang | 2
MSC7RM-CPU-020 | 3
Cuvette temperature high | 3
MSC7RM-CPU-022 | 2
IR source temperature high | 2
MSC7RM-CPU-026 | 2
Homogeniser temperature high | 2
MSC7RM-CPU-027 | 3
Scan length error | 3
MSC7RM-CPU-030 | 2
Mirror movement error | 2
MSC7RM-CPU-031 | 40
IR peak lost | 40
MSC7RM-CPU-033 | 15
Erreur sens miroir | 12
Mirror direction error | 3
MSC7RM-CPU-034 | 4
Crête IR haute | 3
Высокий пик ИК-излучения | 1
MSC7RM-CPU-080 | 109
Pressure stroke power low | 109
MSC7RM-CPU-081 | 100
Bubbles detected in rinze/zero | 100
MSC7RM-CPU-085 | 1
No setup data | 1
PC-042-003 | 1
An error found during a Zero | 1
MilkoScan 7 Disconnect | 5
PC-014-112 | 5
Device ping error | 5
Operator | 50
PC-020-001 | 44
Could not create automatic job for sample ID = . | 34
Could not create automatic job for sample ID = 1. | 7
Work session limit size exceeded. | 3
PC-021-008 | 6
Missing ID-Mask | 6
Operator fault | 70
PC-021-000 | 5
Analyser(s) not active | 5
PC-021-010 | 27
Check-sample(s) missing at end of job | 27
PC-021-012 | 2
No bottle in retest position | 2
PC-044-000 | 7
Limits checking is disabled or no limits are set for any prediction models in the product &apos;Payment Testing&apos;. Check cannot be performed. | 6
Limits checking is disabled or no limits are set for any prediction models in the product 'MSC7RM + FM7DC'. Check cannot be performed. | 1
PC-044-002 | 29
Check-sample product is not defined for a job #1 in a product &apos;COMBI-QC&apos; | 3
Check-sample product is not defined for a job #1 in a product &apos;Payment Testing&apos; | 23
Check-sample product is not defined for a job #3 in a product MSC7RM + FM7DC» | 3
Operator Liquids | 3119
FM7DC-IC-020 | 5
Dye supply almost empty. Halting measurements. | 5
FM7DC-MFH-024 | 4
No dye available. | 4
FM7DC-RSH-001 | 2
Rinse and sheath liquid chamber empty. | 2
FM7-IC-020 | 18
Dye supply almost empty. Halting measurements. | 18
FM7-MFH-024 | 17
No dye available. | 17
FM7-MWH-007 | 27
Internal waste tank sensor missmatch. | 27
FM7-RSH-001 | 72
Rinse and sheath liquid chamber empty. | 72
MSC7RM-CPU-007 | 103
Waste container full | 103
MSC7RM-CPU-059 | 18
Rinse container empty | 18
MSC7RM-CPU-060 | 31
Zero container empty | 31
MSC7RM-CPU-067 | 2822
Kein Abfallbehälter? | 1
No waste container? | 1584
No waste container? Waste container | 1237
Sample errors | 1176
PC-019-002 | 152
Job: &apos;Check-sample definition 1 (1)&apos; reported errors | 1
Job: &apos;Clean x5&apos; reported errors | 12
Job: &apos;DEFINITION LAIT DE CONTROLE A3 (9)&apos; erreur relevées | 1
Job: &apos;Fossomatic 7 adjustment Check (9)&apos; reported errors | 1
Job: &apos;NMR FM Cells Repeat x1&apos; reported errors | 5
Job: &apos;Repeatability (1)&apos; hat Fehler gemeldet | 129
Job: «Определение эталонной пробы 1 (1)» reported errors | 3
PC-042-002 | 49
An error found during a Zero | 49
PC-042-003 | 19
An error found during a Zero | 19
PC-044-006 | 933
Deviation is outside limit. | 933
PC-044-009 | 11
Mean deviation is outside limit. | 11
PC-044-666 | 12
Empty Check Sample | 12
Secondary error | 47
FM7DC-MFH-021 | 6
Internal communication error. | 6
PC-021-007 | 41
Sampling could not complete | 41
Software Error | 4
PC-035-001 | 4
Failed to retrieve alarm settings | 4
Software exceptions | 206
PC-004-002 | 204
An unhandled error was caught in the server. The error was caught in the file d:\b\a1\_work\47\s\foss integrator\domain layer\fcsimtserver\comcommand.cpp line 200. | 204
PC-008-009 | 1
Unit Comm Layer Client Interface - Hidden window is not valid. | 1
PC-008-016 | 1
Unhandled, unexpected fatal error | 1
Software failures | 79
PC-013-007 | 77
Failed to make snapshot of system settings. | 77
PC-016-103 | 2
A completed result submitted to the Result Server was blocked by an incomplete result. Forcing completion of the incomplete result. | 2
Software update issues | 41
FM7CM-CPU-100 | 5
Could not make TcpIp connection to Fossomatic 7 Core | 5
PC-014-001 | 1
Unexpected download reply | 1
PC-014-102 | 8
Device initialization failed | 8
PC-014-120 | 9
Device initialization failed | 9
PC-014-133 | 3
Incompatible boot code | 3
PC-014-137 | 1
Timeout during set device address | 1
PC-014-143 | 1
Timeout waiting on device ready status | 1
PC-014-150 | 13
Connect request to device Fossomatic 7 did not succeed | 13
UA Box - USB disconnect | 12
PC-006-002 | 12
Connection to UA-Interface has been lost and cannot be re-established. Please restart system. | 12
Grand Total | 8914

## Event logs from customer
Instrument Count: |  | Date start | Date end | Customer | Workstation | FIA Version | Instrument | Error type | Error code | Count | Description
9 | 1 | 2019.12.01 | 2020.03.02 | AgroPlem | DESKTOP-P849GT1 | 3.1.1 | Conveyor 7 | Conveyor 7 Operator | CV7-CV7-002 | 4 | Rack detected in front of one or more sensors. Please rewind rack
[8 empty] | Conveyor 7 Calibration | CV7-CV7-003 | 15 | Rack size mismatch. Check if settings and sensor calibrations are correct
[8 empty] | Conveyor 7 Calibration | CV7-CV7-004 | 74 | Rack movement error. Check if the movement is blocked
[8 empty] | Conveyor 7 Calibration | CV7-CV7-009 | 704 | Pipette or Stirrer counter error.
[7 empty] | Fossomatic 7 | Fossomatic 7 Disconnect | FM7-IC-010 | 22 | Instrument unexpectedly reset. Sample under pipette might be contaminated.
[8 empty] | Operator Liquids | FM7-IC-020 | 3 | Dye supply almost empty. Halting measurements.
[8 empty] | Operator Liquids | FM7-MFH-024 | 1 | No dye available.
[8 empty] | FM 7 Instrument | FM7-MWH-006 | 24 | Internal waste tank full.
[8 empty] | Operator Liquids | FM7-MWH-007 | 27 | Internal waste tank sensor missmatch.
[8 empty] | Operator Liquids | FM7-RSH-001 | 22 | Rinse and sheath liquid chamber empty.
[8 empty] | FM 7 Instrument | FM7-RSH-011 | 1 | Rinse-Sheath liquid temperature low.
[8 empty] | FM 7 Instrument | FM7-RSH-014 | 2 | Timeout when filling Rinse / Sheath unit.
[7 empty] | MilkoScan 7 | MilkoScan 7 | MSC7RM-CPU-001 | 3 | Accumulator fill timeout
[8 empty] | MilkoScan 7 | MSC7RM-CPU-003 | 2 | Accumulator normal timeout
[8 empty] | MilkoScan 7 | MSC7RM-CPU-031 | 3 | IR peak lost
[8 empty] | Operator Liquids | MSC7RM-CPU-059 | 9 | Rinse container empty
[8 empty] | Operator Liquids | MSC7RM-CPU-060 | 6 | Zero container empty
[8 empty] | MilkoScan 7 | MSC7RM-CPU-080 | 9 | Pressure stroke power low
[8 empty] | MilkoScan 7 | MSC7RM-CPU-081 | 6 | Bubbles detected in rinze/zero
[7 empty] | PC | Software exceptions | PC-004-002 | 63 | An unhandled error was caught in the server. The error was caught in the file d:\b\a1\_work\47\s\foss integrator\domain layer\fcsimtserver\comcommand.cpp line 200.
[8 empty] | Fossomatic 7 Disconnect | PC-006-141 | 31 | Instrument communication error
[8 empty] | Software failures | PC-013-007 | 41 | Failed to make snapshot of system settings.
[8 empty] | Software update issues | PC-014-102 | 1 | Device initialization failed
[8 empty] | Fossomatic 7 Disconnect | PC-014-119 | 12 | Device failed during device mode change
[8 empty] | Conveyor 7 Disconnect | PC-014-127 | 39 | Connection to Conveyor 7 is lost.
[8 empty] | Software update issues | PC-014-137 | 1 | Timeout during set device address
[8 empty] | Software update issues | PC-014-143 | 1 | Timeout waiting on device ready status
[8 empty] | Software update issues | PC-014-150 | 1 | Connect request to device Fossomatic 7 did not succeed
[8 empty] | Secondary error | PC-021-007 | 4 | Sampling could not complete
[8 empty] | Conveyor 7 Other issues | PC-021-013 | 1 | Conveyor not rewound before restart
[8 empty] | Sample errors | PC-044-006 | 3 | Deviation is outside limit.
[8 empty] | Sample errors | PC-044-666 | 1 | Empty Check Sample
[8 empty] | Check our time-out - should we extend time-out | PC-049-000 | 4 | Rinse job timed out
[8 empty] | Fossomatic 7 Baseline | PC-120-111 | 32 | Baseline out of limit.
| 1 | 2019.12.01 | 2020.03.02 | AgroPlem | DESKTOP-5G5A37P | 3.1.1 | Conveyor 7 | Conveyor 7 Installation | CV7-CV7-001 | 4 | Setup data are invalid. Check if settings are correct
[8 empty] | Conveyor 7 Operator | CV7-CV7-002 | 16 | Rack detected in front of one or more sensors. Please rewind rack
[8 empty] | Conveyor 7 Calibration | CV7-CV7-003 | 5 | Rack size mismatch. Check if settings and sensor calibrations are correct
[8 empty] | Conveyor 7 Calibration | CV7-CV7-004 | 12 | Rack movement error. Check if the movement is blocked
[8 empty] | Conveyor 7 Calibration | CV7-CV7-009 | 93 | Pipette or Stirrer counter error.
[7 empty] | Fossomatic 7 DC | Fossomatic 7 Disconnect | FM7-IC-010 | 27 | Instrument unexpectedly reset. Sample under pipette might be contaminated.
[8 empty] | Operator Liquids | FM7-IC-020 | 5 | Dye supply almost empty. Halting measurements.
[8 empty] | Operator Liquids | FM7-MFH-024 | 4 | No dye available.
[8 empty] | FM 7 Instrument | FM7-MWH-003 | 1 | Mix syringe motor error.
[8 empty] | FM 7 Instrument | FM7-MWH-006 | 19 | Internal waste tank full.
[8 empty] | FM 7 Instrument | FM7-MWH-007 | 18 | Internal waste tank sensor missmatch.
[8 empty] | Operator Liquids | FM7-RSH-001 | 36 | Rinse and sheath liquid chamber empty.
[8 empty] | FM 7 Instrument | FM7-RSH-011 | 1 | Rinse-Sheath liquid temperature low.
[8 empty] | FM 7 Instrument | FM7-RSH-014 | 5 | Timeout when filling Rinse / Sheath unit.
[8 empty] | FM 7 Instrument | FM7-RSH-018 | 1 | Rinse-Sheath valve 2 driver error.
[7 empty] | MilkoScan 7 | MilkoScan 7 | MSC7RM-CPU-001 | 2 | Accumulator fill timeout
[8 empty] | MilkoScan 7 | MSC7RM-CPU-002 | 30 | Accumulator leak
[8 empty] | MilkoScan 7 | MSC7RM-CPU-003 | 38 | Accumulator normal timeout
[8 empty] | MilkoScan 7 | MSC7RM-CPU-020 | 2 | Cuvette temperature high
[8 empty] | MilkoScan 7 | MSC7RM-CPU-030 | 2 | Mirror movement error
[8 empty] | MilkoScan 7 | MSC7RM-CPU-031 | 7 | IR peak lost
[8 empty] | MilkoScan 7 | MSC7RM-CPU-034 | 1 | Высокий пик ИК-излучения
[8 empty] | Operator Liquids | MSC7RM-CPU-059 | 4 | Rinse container empty
[8 empty] | Operator Liquids | MSC7RM-CPU-060 | 2 | Zero container empty
[8 empty] | MilkoScan 7 | MSC7RM-CPU-080 | 16 | Pressure stroke power low
[8 empty] | MilkoScan 7 | MSC7RM-CPU-081 | 48 | Bubbles detected in rinze/zero
[8 empty] | MilkoScan 7 | MSC7RM-CPU-085 | 1 | No setup data
[7 empty] | PC | Software exceptions | PC-004-002 | 125 | An unhandled error was caught in the server. The error was caught in the file d:\b\a1\_work\47\s\foss integrator\domain layer\fcsimtserver\comcommand.cpp line 200.
[8 empty] | Fossomatic 7 Disconnect | PC-006-141 | 38 | Instrument communication error
[8 empty] | Software failures | PC-013-007 | 10 | Failed to make snapshot of system settings.
[8 empty] | Fossomatic 7 Disconnect | PC-014-118 | 2 | Device failed during device mode change (Standby)
[8 empty] | Fossomatic 7 Disconnect | PC-014-119 | 4 | Device failed during device mode change (Auto)
[8 empty] | Fossomatic 7 Disconnect | PC-014-124 | 2 | Device unexpectedly been reset
[8 empty] | Conveyor 7 Disconnect | PC-014-127 | 41 | Connection to Conveyor 7 is lost.
[8 empty] | Software update issues | PC-014-150 | 4 | Connect request to device Fossomatic 7 did not succeed
[8 empty] | Sample errors | PC-019-002 | 1 | Job: &apos;Fossomatic 7 adjustment Check (9)&apos; reported errors
[8 empty] | Sample errors | PC-042-002 | 49 | An error found during a Zero
[8 empty] | Sample errors | PC-042-003 | 19 | An error found during a Zero
[8 empty] | Sample errors | PC-044-006 | 5 | Deviation is outside limit.
[8 empty] | Sample errors | PC-044-009 | 7 | Mean deviation is outside limit.
[8 empty] | Check our time-out - should we extend time-out | PC-049-000 | 2 | Rinse job timed out
[8 empty] | Fossomatic 7 Baseline | PC-120-111 | 84 | Baseline out of limit.
| 1 | 2019.11.19 | 2020.03.02 | Combofoss s/n 91840951 | DESKTOP-3KTQGPM | 3.1.0 | Conveyor 7 | Conveyor 7 Calibration | CV7-CV7-003 | 42 | Rack size mismatch. Check if settings and sensor calibrations are correct
[8 empty] | Conveyor 7 Calibration | CV7-CV7-004 | 2 | Rack movement error. Check if the movement is blocked
[7 empty] | Fossomatic 7 DC | Fossomatic 7 Disconnect | FM7DC-IC-010 | 2 | Instrument unexpectedly reset. Sample under pipette might be contaminated.
[8 empty] | Cadance 100 issue - solved | FM7DC-MFH-008 | 6 | Internal error waiting for pipette.
[8 empty] | Secondary error | FM7DC-MFH-021 | 6 | Internal communication error.
[8 empty] | Operator Liquids | FM7DC-RSH-001 | 2 | Rinse and sheath liquid chamber empty.
[8 empty] | FM 7 DC Instrument | FM7DC-RSH-011 | 1 | Rinse-Sheath liquid temperature low.
[7 empty] | MilkoScan 7 | MilkoScan 7 | MSC7RM-CPU-031 | 1 | IR peak lost
[8 empty] | Operator Liquids | MSC7RM-CPU-059 | 1 | Rinse container empty
[8 empty] | Operator Liquids | MSC7RM-CPU-060 | 1 | Zero container empty
[8 empty] | MilkoScan 7 | MSC7RM-CPU-080 | 1 | Pressure stroke power low
[8 empty] | MilkoScan 7 | MSC7RM-CPU-081 | 2 | Bubbles detected in rinze/zero
[7 empty] | PC | Software exceptions | PC-004-002 | 16 | An unhandled error was caught in the server. The error was caught in the file d:\b\a1\_work\47\s\foss integrator\domain layer\fcsimtserver\comcommand.cpp line 200.
[8 empty] | Fossomatic 7 Disconnect | PC-006-141 | 2 | Instrument communication error
[8 empty] | Software exceptions | PC-008-009 | 1 | Unit Comm Layer Client Interface - Hidden window is not valid.
[8 empty] | Software exceptions | PC-008-016 | 1 | Unhandled, unexpected fatal error
[8 empty] | Software update issues | PC-014-001 | 1 | Unexpected download reply
[8 empty] | Software update issues | PC-014-102 | 1 | Device initialization failed
[8 empty] | Fossomatic 7 Disconnect | PC-014-124 | 6 | Device unexpectedly been reset
[8 empty] | Conveyor 7 Disconnect | PC-014-127 | 4 | Connection to Conveyor 7 is lost.
[8 empty] | Software update issues | PC-014-150 | 1 | Connect request to device Fossomatic 7 did not succeed
[8 empty] | Sample errors | PC-019-002 | 3 | Job: «Определение эталонной пробы 1 (1)» reported errors
[8 empty] | Operator fault | PC-021-012 | 2 | No bottle in retest position
[8 empty] | Conveyor 7 Other issues | PC-021-013 | 7 | Conveyor not rewound before restart
[8 empty] | Software Error | PC-035-001 | 4 | Failed to retrieve alarm settings
[8 empty] | Operator fault | PC-044-000 | 1 | Limits checking is disabled or no limits are set for any prediction models in the product 'MSC7RM + FM7DC'. Check cannot be performed.
[8 empty] | Operator fault | PC-044-002 | 3 | Check-sample product is not defined for a job #3 in a product MSC7RM + FM7DC»
[8 empty] | Sample errors | PC-044-006 | 15 | Deviation is outside limit.
[8 empty] | Sample errors | PC-044-009 | 4 | Mean deviation is outside limit.
[8 empty] | Check our time-out - should we extend time-out | PC-049-000 | 1 | Rinse job timed out
[8 empty] | Installation | PC-120-108 | 1 | Instrument Configuration must be carried out.
[8 empty] | Flow system Error | PC-120-110 | 4 | Flow system error. Please begin by checking incubation valves. Other possible reasons could be valve VPL_V9, also check pipette for blockage.
[8 empty] | FM 7 Instrument | PC-120-112 | 2 | Malfunctioning pinch valves
| 1 | 2019.12.03 | 2020.03.03 | Lacteus, Gaétan Renard | FOSSPC-France | 2.3.8 | Conveyor 4000 | Conveyor 4000 | CV-CPU-018 | 10 | Rack en face des capteurs. Faire un retour du rack
[8 empty] | Conveyor 4000 | CV-CPU-022 | 6 | Mouvement de base rack
[8 empty] | Conveyor 4000 | CV-CPU-050 | 37 | La taille du rack ne correspond pas
[7 empty] | Fossomatic 7 | Fossomatic 7 Disconnect | FM7-IC-010 | 5 | Instrument unexpectedly reset. Sample under pipette might be contaminated.
[8 empty] | Operator Liquids | FM7-IC-020 | 3 | Dye supply almost empty. Halting measurements.
[8 empty] | FM 7 Instrument | FM7-MFH-007 | 2 | Internal error waiting for sample intake.
[8 empty] | FM 7 Instrument | FM7-MFH-022 | 51 | Sample intake failure (VPL_S1).
[8 empty] | Operator Liquids | FM7-MFH-024 | 2 | No dye available.
[8 empty] | FM 7 Instrument | FM7-MWH-008 | 4 | Mix liquid sensor calibration error. Check supply.
[8 empty] | FM 7 Instrument | FM7-RSH-005 | 1 | inse-Sheath pressure high
[8 empty] | FM 7 Instrument | FM7-RSH-006 | 75 | Rinse-Sheath pressure low
[8 empty] | FM 7 Instrument | FM7-RSH-011 | 7 | Rinse-Sheath liquid temperature low.
[8 empty] | FM 7 Instrument | FM7-RSH-014 | 7 | Timeout when filling Rinse / Sheath unit.
[7 empty] | MilkoScan 7 | MilkoScan 7 | MSC7RM-CPU-001 | 7 | Accumulator fill timeout
[8 empty] | MilkoScan 7 | MSC7RM-CPU-002 | 1 | Accumulator leak
[8 empty] | MilkoScan 7 | MSC7RM-CPU-031 | 1 | IR peak lost
[8 empty] | MilkoScan 7 | MSC7RM-CPU-033 | 12 | Erreur sens miroir
[8 empty] | MilkoScan 7 | MSC7RM-CPU-034 | 3 | Crête IR haute
[8 empty] | Operator Liquids | MSC7RM-CPU-059 | 1 | Rinse container empty
[8 empty] | Operator Liquids | MSC7RM-CPU-060 | 14 | Zero container empty
[8 empty] | MilkoScan 7 | MSC7RM-CPU-080 | 6 | Pressure stroke power low
[8 empty] | MilkoScan 7 | MSC7RM-CPU-081 | 3 | Bubbles detected in rinze/zero
[7 empty] | PC | Fossomatic 7 Disconnect | PC-006-141 | 9 | Instrument communication error
[8 empty] | Software update issues | PC-014-102 | 1 | Device initialization failed
[8 empty] | Ask Ole | PC-014-113 | 2 | No active analyser
[8 empty] | Conveyor 7 Disconnect | PC-014-127 | 17 | Connection to Conveyor 7 is lost.
[8 empty] | Sample errors | PC-019-002 | 1 | Job: &apos;DEFINITION LAIT DE CONTROLE A3 (9)&apos; erreur relevées
[8 empty] | Operator | PC-020-001 | 3 | Work session limit size exceeded.
[8 empty] | Secondary error | PC-021-007 | 3 | Sampling could not complete
[8 empty] | Ask Ole | PC-032-101 | 2 | Read Error Correction dialog synchronisation error
[8 empty] | MilkoScan 7 | PC-042-003 | 1 | An error found during a Zero
[8 empty] | Sample errors | PC-044-006 | 4 | Deviation is outside limit.
[8 empty] | Check our time-out - should we extend time-out | PC-049-000 | 5 | Rinse job timed out
[8 empty] | LIMS Server Error | PC-059-101 | 2 | External data communication server has closed the network connection
| 1 | 2019.12.01 | 2020.03.05 | Waldsieversdorf | FOSSPC | 3.1.1 | Conveyor 7 | Conveyor 7 Operator | CV7-CV7-002 | 3 | Rack detected in front of one or more sensors. Please rewind rack
[8 empty] | Conveyor 7 Calibration | CV7-CV7-003 | 4 | Rack size mismatch. Check if settings and sensor calibrations are correct
[8 empty] | Conveyor 7 Calibration | CV7-CV7-009 | 42 | Pipette or Stirrer counter error.
[7 empty] | Fossomatic 7 DC | Software update issues | FM7CM-CPU-100 | 4 | Could not make TcpIp connection to Fossomatic 7 Core
[8 empty] | FM 7 Instrument | FM7DC-FCU-001 | 423 | Needle inlet temperature too low
[8 empty] | FM 7 Instrument | FM7DC-IC-005 | 7 | Intake bubble sensor (VPL_S1) calibration error.
[8 empty] | Fossomatic 7 Disconnect | FM7DC-IC-010 | 18 | Instrument unexpectedly reset. Sample under pipette might be contaminated.
[8 empty] | Operator Liquids | FM7DC-IC-020 | 5 | Dye supply almost empty. Halting measurements.
[8 empty] | FM 7 Instrument | FM7DC-MFH-002 | 29 | SSY motor timeout, sensor not found.
[8 empty] | FM 7 Instrument | FM7DC-MFH-007 | 36 | Internal error waiting for sample intake.
[8 empty] | FM 7 Instrument | FM7DC-MFH-022 | 86 | Sample intake failure (VPL_S1).
[8 empty] | Operator Liquids | FM7DC-MFH-024 | 4 | No dye available.
[8 empty] | FM 7 Instrument | FM7DC-MWH-006 | 6 | Internal waste tank full.
[8 empty] | FM 7 Instrument | FM7DC-RSH-011 | 13 | Rinse-Sheath liquid temperature low.
[8 empty] | FM 7 Instrument | FM7DC-RSH-014 | 6 | Timeout when filling Rinse / Sheath unit.
[7 empty] | MilkoScan 7 | MilkoScan 7 | MSC7RM-CPU-001 | 8 | Accumulator fill timeout
[8 empty] | MilkoScan 7 | MSC7RM-CPU-009 | 2 | Messdauer zu lang
[8 empty] | MilkoScan 7 | MSC7RM-CPU-031 | 20 | IR peak lost
[8 empty] | Operator Liquids | MSC7RM-CPU-059 | 3 | Rinse container empty
[8 empty] | Operator Liquids | MSC7RM-CPU-060 | 8 | Zero container empty
[8 empty] | Operator Liquids | MSC7RM-CPU-067 | 1 | Kein Abfallbehälter?
[8 empty] | MilkoScan 7 | MSC7RM-CPU-080 | 28 | Pressure stroke power low
[8 empty] | MilkoScan 7 | MSC7RM-CPU-081 | 6 | Bubbles detected in rinze/zero
[7 empty] | PC | Fossomatic 7 Disconnect | PC-006-141 | 8 | Instrument communication error
[8 empty] | Software failures | PC-013-007 | 26 | Failed to make snapshot of system settings.
[8 empty] | Fossomatic 7 Disconnect | PC-014-118 | 2 | Device failed during device mode change (Standby)
[8 empty] | Conveyor 7 Disconnect | PC-014-127 | 9 | Connection to Conveyor 7 is lost.
[8 empty] | Software update issues | PC-014-150 | 7 | Connect request to device Fossomatic 7 did not succeed
[8 empty] | Software failures | PC-016-103 | 1 | A completed result submitted to the Result Server was blocked by an incomplete result. Forcing completion of the incomplete result.
[8 empty] | Sample errors | PC-019-002 | 129 | Job: &apos;Repeatability (1)&apos; hat Fehler gemeldet
[8 empty] | Operator fault | PC-021-000 | 2 | Analyser(s) not active
[8 empty] | Secondary error | PC-021-007 | 29 | Sampling could not complete
[8 empty] | Ask Ole | PC-021-009 | 2 | Job type error
[8 empty] | Operator fault | PC-021-010 | 27 | Check-sample(s) missing at end of job
[8 empty] | Conveyor 7 Other issues | PC-021-013 | 3 | Conveyor not rewound before restart
[8 empty] | Sample errors | PC-044-006 | 198 | Deviation is outside limit.
[8 empty] | Check our time-out - should we extend time-out | PC-049-000 | 2 | Rinse job timed out
[8 empty] | LIMS Server Error | PC-059-101 | 12 | External data communication server has closed the network connection
[8 empty] | Flow system Error | PC-120-109 | 1 | Baseline level is saturated. No or too little solution is passing through the flow cytometer.
[8 empty] | Flow system Error | PC-120-110 | 3 | Flow system error. Please begin by checking incubation valves. Other possible reasons could be valve VPL_V9, also check pipette for blockage.
| 1 | 2019.12.01 | 2020.03.03 | Elita (Kazan) |  |  | Conveyor 7 | Conveyor 7 Calibration | CV7-CV7-003 | 2 | Rack size mismatch. Check if settings and sensor calibrations are correct
[8 empty] | Conveyor 7 Calibration | CV7-CV7-004 | 1 | Rack movement error. Check if the movement is blocked
[8 empty] | Conveyor 7 Calibration | CV7-CV7-009 | 30 | Pipette or Stirrer counter error.
[7 empty] | Fossomatic 7 | Software update issues | FM7CM-CPU-100 | 1 | Could not make TcpIp connection to Fossomatic 7 Core
[8 empty] | Fossomatic 7 Disconnect | FM7-IC-010 | 7 | Instrument unexpectedly reset. Sample under pipette might be contaminated.
[8 empty] | Fossomatic 7 Disconnect | FM7-MFH-020 | 1 | Internal communication error.
[8 empty] | FM 7 Instrument | FM7-MWH-006 | 114 | Internal waste tank full.
[8 empty] | FM 7 Instrument | FM7-MWH-007 | 120 | Internal waste tank sensor missmatch.
[8 empty] | Operator Liquids | FM7-RSH-001 | 1 | Rinse and sheath liquid chamber empty.
[8 empty] | FM 7 Instrument | FM7-RSH-011 | 1 | Rinse-Sheath liquid temperature low.
[7 empty] | MilkoScan 7 | MilkoScan 7 | MSC7RM-CPU-001 | 1 | Accumulator fill timeout
[8 empty] | MilkoScan 7 | MSC7RM-CPU-003 | 100 | Accumulator normal timeout
[8 empty] | MilkoScan 7 | MSC7RM-CPU-031 | 2 | IR peak lost
[8 empty] | MilkoScan 7 | MSC7RM-CPU-033 | 1 | Mirror direction error
[8 empty] | MilkoScan 7 | MSC7RM-CPU-080 | 1 | Pressure stroke power low
[8 empty] | MilkoScan 7 | MSC7RM-CPU-081 | 1 | Bubbles detected in rinze/zero
[7 empty] | PC | Fossomatic 7 Disconnect | PC-006-141 | 4 | Instrument communication error
[8 empty] | Software update issues | PC-014-102 | 1 | Device initialization failed
[8 empty] | MilkoScan 7 Disconnect | PC-014-112 | 1 | Device ping error
[8 empty] | Fossomatic 7 Disconnect | PC-014-119 | 10 | Device failed during device mode change
[8 empty] | Fossomatic 7 Disconnect | PC-014-124 | 6 | Device unexpectedly been reset
[8 empty] | Conveyor 7 Disconnect | PC-014-127 | 5 | Connection to Conveyor 7 is lost.
[8 empty] | Operator fault | PC-021-000 | 3 | Analyser(s) not active
[8 empty] | Ask Ole | PC-021-009 | 3 | Job type error
[8 empty] | Check our time-out - should we extend time-out | PC-049-000 | 2 | Rinse job timed out
[8 empty] | LIMS Server Error | PC-059-101 | 1 | External data communication server has closed the network connection
[8 empty] | Fossomatic 7 Baseline | PC-120-111 | 9 | Baseline out of limit.
| 1 | 2019.12.31 | 2020.03.31 | NML | DESKTOP-3JI1G3S | 3.0.0 | Conveyor 7 | Conveyor 7 Operator | CV7-CV7-002 | 1 | Rack detected in front of one or more sensors. Please rewind rack
[8 empty] | Conveyor 7 Calibration | CV7-CV7-004 | 2 | Rack movement error. Check if the movement is blocked
[7 empty] | Fossomatic 7 | FM 7 Instrument | FM7-IC-005 | 63 | Intake bubble sensor (VPL_S1) calibration error.
[8 empty] | Fossomatic 7 Disconnect | FM7-IC-010 | 4 | Instrument unexpectedly reset. Sample under pipette might be contaminated.
[8 empty] | FM 7 Instrument | FM7-MWH-006 | 22 | Internal waste tank full.
[8 empty] | Operator Liquids | FM7-RSH-001 | 5 | Rinse and sheath liquid chamber empty.
[8 empty] | FM 7 Instrument | FM7-RSH-014 | 3 | Timeout when filling Rinse / Sheath unit.
[7 empty] | MilkoScan 7 | MilkoScan 7 | MSC7RM-CPU-001 | 12 | Accumulator fill timeout
[8 empty] | MilkoScan 7 | MSC7RM-CPU-002 | 4 | Accumulator leak
[8 empty] | MilkoScan 7 | MSC7RM-CPU-003 | 9 | Accumulator normal timeout
[8 empty] | Operator Liquids | MSC7RM-CPU-007 | 83 | Waste container full
[8 empty] | MilkoScan 7 | MSC7RM-CPU-009 | 5 | Measure time too long
[8 empty] | MilkoScan 7 | MSC7RM-CPU-020 | 1 | Cuvette temperature high
[8 empty] | Operator Liquids | MSC7RM-CPU-067 | 276 | No waste container?
[8 empty] | MilkoScan 7 | MSC7RM-CPU-080 | 16 | Pressure stroke power low
[8 empty] | MilkoScan 7 | MSC7RM-CPU-081 | 34 | Bubbles detected in rinze/zero
[7 empty] | PC | Fossomatic 7 Disconnect | PC-006-141 | 5 | Instrument communication error
[8 empty] | MilkoScan 7 Disconnect | PC-014-112 | 1 | Device ping error
[8 empty] | Fossomatic 7 Disconnect | PC-014-117 | 1 | Device failed during device mode change
[8 empty] | Fossomatic 7 Disconnect | PC-014-124 | 2 | Device unexpectedly been reset
[8 empty] | Fossomatic 7 Disconnect | PC-014-127 | 6 | Connection to Fossomatic 7 is lost.
[8 empty] | Sample errors | PC-019-002 | 1 | Job: &apos;Check-sample definition 1 (1)&apos; reported errors
[8 empty] | Operator | PC-020-001 | 7 | Could not create automatic job for sample ID = 1.
[8 empty] | Operator | PC-021-008 | 2 | Missing ID-Mask
[8 empty] | Ask Ole | PC-032-101 | 4 | Read Error Correction dialog synchronisation error
[8 empty] | Sample errors | PC-044-006 | 86 | Deviation is outside limit.
| 1 | 2019.12.31 | 2020.03.31 | NML | NMR14261 | 2.3.8 | Conveyor 7 | Conveyor 7 Calibration | CV7-CV7-003 | 3 | Rack size mismatch. Check if settings are correct
[8 empty] | Conveyor 7 Calibration | CV7-CV7-004 | 12 | Rack movement error. Check if the movement is blocked
[7 empty] | Fossomatic 7 | FM 7 Instrument | FM7-FCU-003 | 11 | Laser detector receives too low signal from laser
[8 empty] | FM 7 Instrument | FM7-IC-001 | 1 | Incubation module not found at startup.
[8 empty] | FM 7 Instrument | FM7-IC-005 | 3 | Intake bubble sensor (VPL_S1) calibration error.
[8 empty] | Fossomatic 7 Disconnect | FM7-IC-010 | 14 | Instrument unexpectedly reset. Sample under pipette might be contaminated.
[8 empty] | Operator Liquids | FM7-IC-020 | 2 | Dye supply almost empty. Halting measurements.
[8 empty] | Operator Liquids | FM7-MFH-024 | 4 | No dye available.
[8 empty] | FM 7 Instrument | FM7-MFH-025 | 1 | Intake heater temperature low.
[8 empty] | FM 7 Instrument | FM7-RSH-014 | 2 | Timeout when filling Rinse / Sheath unit.
[7 empty] | MilkoScan 7 | MilkoScan 7 | MSC7RM-CPU-001 | 70 | Accumulator fill timeout
[8 empty] | MilkoScan 7 | MSC7RM-CPU-002 | 16 | Accumulator leak
[8 empty] | MilkoScan 7 | MSC7RM-CPU-003 | 5 | Accumulator normal timeout
[8 empty] | Operator Liquids | MSC7RM-CPU-007 | 15 | Waste container full
[8 empty] | MilkoScan 7 | MSC7RM-CPU-009 | 69 | Measure time too long
[8 empty] | MilkoScan 7 | MSC7RM-CPU-022 | 2 | IR source temperature high
[8 empty] | MilkoScan 7 | MSC7RM-CPU-026 | 2 | Homogeniser temperature high
[8 empty] | MilkoScan 7 | MSC7RM-CPU-027 | 1 | Scan length error
[8 empty] | MilkoScan 7 | MSC7RM-CPU-031 | 3 | IR peak lost
[8 empty] | MilkoScan 7 | MSC7RM-CPU-033 | 1 | Mirror direction error
[8 empty] | Operator Liquids | MSC7RM-CPU-067 | 1237 | No waste container? Waste container
[8 empty] | MilkoScan 7 | MSC7RM-CPU-080 | 2 | Pressure stroke power low
[7 empty] | PC | UA Box - USB disconnect | PC-006-002 | 3 | Connection to UA-Interface has been lost and cannot be re-established. Please restart system.
[8 empty] | Fossomatic 7 Disconnect | PC-006-141 | 5 | Instrument communication error
[8 empty] | Ask Ole | PC-013-006 | 4 | Data from unit was discarded
[8 empty] | MilkoScan 7 Disconnect | PC-014-112 | 2 | Device ping error
[8 empty] | Fossomatic 7 Disconnect | PC-014-117 | 3 | Device failed during device mode change
[8 empty] | Software update issues | PC-014-120 | 2 | Device initialization failed
[8 empty] | Fossomatic 7 Disconnect | PC-014-124 | 8 | Device unexpectedly been reset
[8 empty] | Conveyor 7 Disconnect | PC-014-127 | 5 | Connection to Conveyor 7 is lost.
[8 empty] | LIMS Server Error | PC-016-000 | 1 | Export of result transmission has timed out.
[8 empty] | Sample errors | PC-019-002 | 12 | Job: &apos;Clean x5&apos; reported errors
[8 empty] | Operator | PC-020-001 | 16 | Could not create automatic job for sample ID = .
[8 empty] | Operator | PC-021-008 | 4 | Missing ID-Mask
[8 empty] | Conveyor 7 Other issues | PC-021-013 | 2 | Conveyor not rewound before restart
[8 empty] | Ask Ole | PC-032-101 | 11 | Read Error Correction dialog synchronisation error
[8 empty] | Operator fault | PC-044-000 | 6 | Limits checking is disabled or no limits are set for any prediction models in the product &apos;Payment Testing&apos;. Check cannot be performed.
[8 empty] | Operator fault | PC-044-002 | 3 | Check-sample product is not defined for a job #1 in a product &apos;COMBI-QC&apos;
[8 empty] | Sample errors | PC-044-006 | 333 | Deviation is outside limit.
[8 empty] | Check our time-out - should we extend time-out | PC-049-000 | 1 | Rinse job timed out
[8 empty] | LIMS Server Error | PC-059-101 | 88 | External data communication server has closed the network connection
[8 empty] | Installation | PC-071-050 | 4 | Failed to verify that COM port is configured
| 1 | 2019.12.31 | 2020.03.31 | NML | Wol_Combi4 | 2.3.8 | Conveyor | Conveyor 7 Operator | CV7-CV7-002 | 8 | Rack detected in front of one or more sensors. Please rewind rack
[8 empty] | Conveyor 7 Calibration | CV7-CV7-004 | 34 | Rack movement error. Check if the movement is blocked
[7 empty] | Fossomatic 7 | FM 7 Instrument | FM7-FCU-003 | 1 | Laser detector receives too low signal from laser
[8 empty] | FM 7 Instrument | FM7-IC-005 | 163 | Intake bubble sensor (VPL_S1) calibration error.
[8 empty] | Fossomatic 7 Disconnect | FM7-IC-010 | 80 | Instrument unexpectedly reset. Sample under pipette might be contaminated.
[8 empty] | Operator Liquids | FM7-IC-020 | 5 | Dye supply almost empty. Halting measurements.
[8 empty] | FM 7 Instrument | FM7-MFH-022 | 14 | Sample intake failure (VPL_S1).
[8 empty] | Operator Liquids | FM7-MFH-024 | 6 | No dye available.
[8 empty] | Operator Liquids | FM7-RSH-001 | 8 | Rinse and sheath liquid chamber empty.
[8 empty] | FM 7 Instrument | FM7-RSH-006 | 16 | Rinse-Sheath pressure low
[8 empty] | FM 7 Instrument | FM7-RSH-014 | 3 | Timeout when filling Rinse / Sheath unit.
[7 empty] | MilkoScan 7 | MilkoScan 7 | MSC7RM-CPU-001 | 8 | Accumulator fill timeout
[8 empty] | MilkoScan 7 | MSC7RM-CPU-002 | 3 | Accumulator leak
[8 empty] | MilkoScan 7 | MSC7RM-CPU-003 | 10 | Accumulator normal timeout
[8 empty] | Operator Liquids | MSC7RM-CPU-007 | 5 | Waste container full
[8 empty] | MilkoScan 7 | MSC7RM-CPU-008 | 1 | H-pump stroke-timeout
[8 empty] | MilkoScan 7 | MSC7RM-CPU-009 | 5 | Measure time too long
[8 empty] | MilkoScan 7 | MSC7RM-CPU-027 | 2 | Scan length error
[8 empty] | MilkoScan 7 | MSC7RM-CPU-031 | 3 | IR peak lost
[8 empty] | MilkoScan 7 | MSC7RM-CPU-033 | 1 | Mirror direction error
[8 empty] | Operator Liquids | MSC7RM-CPU-067 | 1308 | No waste container?
[8 empty] | MilkoScan 7 | MSC7RM-CPU-080 | 30 | Pressure stroke power low
[7 empty] | PC | UA Box - USB disconnect | PC-006-002 | 9 | Connection to UA-Interface has been lost and cannot be re-established. Please restart system.
[8 empty] | Installation | PC-006-124 | 2 | Cannot Initialize Gateways
[8 empty] | Installation | PC-006-140 | 2 | No hardware driver found
[8 empty] | Fossomatic 7 Disconnect | PC-006-141 | 15 | Instrument communication error
[8 empty] | Ask Ole | PC-013-006 | 1 | Data from unit was discarded
[8 empty] | Ask Ole | PC-013-031 | 1 | Data lost from device
[8 empty] | Software update issues | PC-014-102 | 4 | Device initialization failed
[8 empty] | MilkoScan 7 Disconnect | PC-014-112 | 1 | Device ping error
[8 empty] | Installation | PC-014-115 | 1 | Unable to connect to device(s).
[8 empty] | Software update issues | PC-014-120 | 7 | Device initialization failed
[8 empty] | Conveyor 7 Disconnect | PC-014-127 | 8 | Connection to Conveyor 7 is lost.
[8 empty] | Software update issues | PC-014-133 | 3 | Incompatible boot code
[8 empty] | LIMS Server Error | PC-016-000 | 3 | Export of result transmission has timed out.
[8 empty] | Software failures | PC-016-103 | 1 | A completed result submitted to the Result Server was blocked by an incomplete result. Forcing completion of the incomplete result.
[8 empty] | Sample errors | PC-019-002 | 5 | Job: &apos;NMR FM Cells Repeat x1&apos; reported errors
[8 empty] | Operator | PC-020-001 | 18 | Could not create automatic job for sample ID = .
[8 empty] | Secondary error | PC-021-007 | 5 | Sampling could not complete
[8 empty] | Conveyor 7 Other issues | PC-021-013 | 3 | Conveyor not rewound before restart
[8 empty] | Ask Ole | PC-032-101 | 10 | Read Error Correction dialog synchronisation error
[8 empty] | Operator fault | PC-044-002 | 23 | Check-sample product is not defined for a job #1 in a product &apos;Payment Testing&apos;
[8 empty] | Sample errors | PC-044-006 | 289 | Deviation is outside limit.
[8 empty] | Sample errors | PC-044-666 | 11 | Empty Check Sample
[8 empty] | Check our time-out - should we extend time-out | PC-049-000 | 9 | Rinse job timed out
[8 empty] | LIMS Server Error | PC-059-101 | 51 | External data communication server has closed the network connection
[8 empty] | LIMS Server Error | PC-071-001 | 1 | Communication failure on the CS83 on RS232 (Serial) data connection
[8 empty] | LIMS Server Error | PC-074-201 | 1 | An error occurred during database export

## Grouped Events
Row Labels | Sum of Count
Ask Ole | 70
PC-013-006 | 5
Data from unit was discarded | 5
PC-013-031 | 1
Data lost from device | 1
PC-014-113 | 2
No active analyser | 2
PC-021-009 | 5
Job type error | 5
PC-032-101 | 27
Read Error Correction dialog synchronisation error | 27
PC-035-001 | 4
Failed to retrieve alarm settings | 4
PC-049-000 | 26
Rinse job timed out | 26
Cadance 100 issue - solved | 6
FM7DC-MFH-008 | 6
Internal error waiting for pipette. | 6
Conveyor 4000 | 53
CV-CPU-018 | 10
Rack en face des capteurs. Faire un retour du rack | 10
CV-CPU-022 | 6
Mouvement de base rack | 6
CV-CPU-050 | 37
La taille du rack ne correspond pas | 37
Conveyor 7 Calibration | 1077
CV7-CV7-003 | 71
Rack size mismatch. Check if settings and sensor calibrations are correct | 68
Rack size mismatch. Check if settings are correct | 3
CV7-CV7-004 | 137
Rack movement error. Check if the movement is blocked | 137
CV7-CV7-009 | 869
Pipette or Stirrer counter error. | 869
Conveyor 7 Disconnect | 134
PC-014-127 | 134
Connection to Conveyor 7 is lost. | 128
Connection to Fossomatic 7 is lost. | 6
Conveyor 7 Installation | 4
CV7-CV7-001 | 4
Setup data are invalid. Check if settings are correct | 4
Conveyor 7 Operator | 32
CV7-CV7-002 | 32
Rack detected in front of one or more sensors. Please rewind rack | 32
Conveyor 7 Other issues | 16
PC-021-013 | 16
Conveyor not rewound before restart | 16
Flow system Error | 8
PC-120-109 | 1
Baseline level is saturated. No or too little solution is passing through the flow cytometer. | 1
PC-120-110 | 7
Flow system error. Please begin by checking incubation valves. Other possible reasons could be valve VPL_V9, also check pipette for blockage. | 7
FM 7 DC Instrument | 1
FM7DC-RSH-011 | 1
Rinse-Sheath liquid temperature low. | 1
FM 7 Instrument | 1364
FM7DC-FCU-001 | 423
Needle inlet temperature too low | 423
FM7DC-IC-005 | 7
Intake bubble sensor (VPL_S1) calibration error. | 7
FM7DC-MFH-002 | 29
SSY motor timeout, sensor not found. | 29
FM7DC-MFH-007 | 36
Internal error waiting for sample intake. | 36
FM7DC-MFH-022 | 86
Sample intake failure (VPL_S1). | 86
FM7DC-MWH-006 | 6
Internal waste tank full. | 6
FM7DC-RSH-011 | 13
Rinse-Sheath liquid temperature low. | 13
FM7DC-RSH-014 | 6
Timeout when filling Rinse / Sheath unit. | 6
FM7-FCU-003 | 12
Laser detector receives too low signal from laser | 12
FM7-IC-001 | 1
Incubation module not found at startup. | 1
FM7-IC-005 | 229
Intake bubble sensor (VPL_S1) calibration error. | 229
FM7-MFH-007 | 2
Internal error waiting for sample intake. | 2
FM7-MFH-022 | 65
Sample intake failure (VPL_S1). | 65
FM7-MFH-025 | 1
Intake heater temperature low. | 1
FM7-MWH-003 | 1
Mix syringe motor error. | 1
FM7-MWH-006 | 179
Internal waste tank full. | 179
FM7-MWH-007 | 138
Internal waste tank sensor missmatch. | 138
FM7-MWH-008 | 4
Mix liquid sensor calibration error. Check supply. | 4
FM7-RSH-005 | 1
inse-Sheath pressure high | 1
FM7-RSH-006 | 91
Rinse-Sheath pressure low | 91
FM7-RSH-011 | 10
Rinse-Sheath liquid temperature low. | 10
FM7-RSH-014 | 22
Timeout when filling Rinse / Sheath unit. | 22
PC-120-112 | 2
Malfunctioning pinch valves | 2
Fossomatic 7 Baseline | 125
PC-120-111 | 125
Baseline out of limit. | 125
Fossomatic 7 Disconnect | 355
FM7DC-IC-010 | 20
Instrument unexpectedly reset. Sample under pipette might be contaminated. | 20
FM7-IC-010 | 159
Instrument unexpectedly reset. Sample under pipette might be contaminated. | 159
FM7-MFH-020 | 1
Internal communication error. | 1
PC-006-141 | 117
Instrument communication error | 117
PC-014-117 | 4
Device failed during device mode change | 4
PC-014-118 | 4
Device failed during device mode change (Standby) | 4
PC-014-119 | 26
Device failed during device mode change | 22
Device failed during device mode change (Auto) | 4
PC-014-124 | 24
Device unexpectedly been reset | 24
Hardware failure | 1
FM7-RSH-018 | 1
Rinse-Sheath valve 2 driver error. | 1
Installation | 10
PC-006-124 | 2
Cannot Initialize Gateways | 2
PC-006-140 | 2
No hardware driver found | 2
PC-014-115 | 1
Unable to connect to device(s). | 1
PC-071-050 | 4
Failed to verify that COM port is configured | 4
PC-120-108 | 1
Instrument Configuration must be carried out. | 1
LIMS Server Error | 160
PC-016-000 | 4
Export of result transmission has timed out. | 4
PC-059-101 | 154
External data communication server has closed the network connection | 154
PC-071-001 | 1
Communication failure on the CS83 on RS232 (Serial) data connection | 1
PC-074-201 | 1
An error occurred during database export | 1
MilkoScan 7 | 694
MSC7RM-CPU-001 | 111
Accumulator fill timeout | 111
MSC7RM-CPU-002 | 54
Accumulator leak | 54
MSC7RM-CPU-003 | 164
Accumulator normal timeout | 164
MSC7RM-CPU-008 | 1
H-pump stroke-timeout | 1
MSC7RM-CPU-009 | 81
Measure time too long | 79
Messdauer zu lang | 2
MSC7RM-CPU-020 | 3
Cuvette temperature high | 3
MSC7RM-CPU-022 | 2
IR source temperature high | 2
MSC7RM-CPU-026 | 2
Homogeniser temperature high | 2
MSC7RM-CPU-027 | 3
Scan length error | 3
MSC7RM-CPU-030 | 2
Mirror movement error | 2
MSC7RM-CPU-031 | 40
IR peak lost | 40
MSC7RM-CPU-033 | 15
Erreur sens miroir | 12
Mirror direction error | 3
MSC7RM-CPU-034 | 4
Crête IR haute | 3
Высокий пик ИК-излучения | 1
MSC7RM-CPU-067 | 1
Kein Abfallbehälter? | 1
MSC7RM-CPU-080 | 109
Pressure stroke power low | 109
MSC7RM-CPU-081 | 100
Bubbles detected in rinze/zero | 100
MSC7RM-CPU-085 | 1
No setup data | 1
PC-042-003 | 1
An error found during a Zero | 1
MilkoScan 7 Disconnect | 5
PC-014-112 | 5
Device ping error | 5
Operator | 50
PC-020-001 | 44
Could not create automatic job for sample ID = . | 34
Could not create automatic job for sample ID = 1. | 7
Work session limit size exceeded. | 3
PC-021-008 | 6
Missing ID-Mask | 6
Operator fault | 70
PC-021-000 | 5
Analyser(s) not active | 5
PC-021-010 | 27
Check-sample(s) missing at end of job | 27
PC-021-012 | 2
No bottle in retest position | 2
PC-044-000 | 7
Limits checking is disabled or no limits are set for any prediction models in the product &apos;Payment Testing&apos;. Check cannot be performed. | 6
Limits checking is disabled or no limits are set for any prediction models in the product 'MSC7RM + FM7DC'. Check cannot be performed. | 1
PC-044-002 | 29
Check-sample product is not defined for a job #1 in a product &apos;COMBI-QC&apos; | 3
Check-sample product is not defined for a job #1 in a product &apos;Payment Testing&apos; | 23
Check-sample product is not defined for a job #3 in a product MSC7RM + FM7DC» | 3
Operator Liquids | 3118
FM7DC-IC-020 | 5
Dye supply almost empty. Halting measurements. | 5
FM7DC-MFH-024 | 4
No dye available. | 4
FM7DC-RSH-001 | 2
Rinse and sheath liquid chamber empty. | 2
FM7-IC-020 | 18
Dye supply almost empty. Halting measurements. | 18
FM7-MFH-024 | 17
No dye available. | 17
FM7-MWH-007 | 27
Internal waste tank sensor missmatch. | 27
FM7-RSH-001 | 72
Rinse and sheath liquid chamber empty. | 72
MSC7RM-CPU-007 | 103
Waste container full | 103
MSC7RM-CPU-059 | 18
Rinse container empty | 18
MSC7RM-CPU-060 | 31
Zero container empty | 31
MSC7RM-CPU-067 | 2821
No waste container? | 2821
Sample errors | 1176
PC-019-002 | 152
Job: &apos;Check-sample definition 1 (1)&apos; reported errors | 1
Job: &apos;Clean x5&apos; reported errors | 12
Job: &apos;DEFINITION LAIT DE CONTROLE A3 (9)&apos; erreur relevées | 1
Job: &apos;Fossomatic 7 adjustment Check (9)&apos; reported errors | 1
Job: &apos;NMR FM Cells Repeat x1&apos; reported errors | 5
Job: &apos;Repeatability (1)&apos; hat Fehler gemeldet | 129
Job: «Определение эталонной пробы 1 (1)» reported errors | 3
PC-042-002 | 49
An error found during a Zero | 49
PC-042-003 | 19
An error found during a Zero | 19
PC-044-006 | 933
Deviation is outside limit. | 933
PC-044-009 | 11
Mean deviation is outside limit. | 11
PC-044-666 | 12
Empty Check Sample | 12
Secondary error | 47
FM7DC-MFH-021 | 6
Internal communication error. | 6
PC-021-007 | 41
Sampling could not complete | 41
Software exceptions | 206
PC-004-002 | 204
An unhandled error was caught in the server. The error was caught in the file d:\b\a1\_work\47\s\foss integrator\domain layer\fcsimtserver\comcommand.cpp line 200. | 204
PC-008-009 | 1
Unit Comm Layer Client Interface - Hidden window is not valid. | 1
PC-008-016 | 1
Unhandled, unexpected fatal error | 1
Software failures | 79
PC-013-007 | 77
Failed to make snapshot of system settings. | 77
PC-016-103 | 2
A completed result submitted to the Result Server was blocked by an incomplete result. Forcing completion of the incomplete result. | 2
Software update issues | 41
FM7CM-CPU-100 | 5
Could not make TcpIp connection to Fossomatic 7 Core | 5
PC-014-001 | 1
Unexpected download reply | 1
PC-014-102 | 8
Device initialization failed | 8
PC-014-120 | 9
Device initialization failed | 9
PC-014-133 | 3
Incompatible boot code | 3
PC-014-137 | 1
Timeout during set device address | 1
PC-014-143 | 1
Timeout waiting on device ready status | 1
PC-014-150 | 13
Connect request to device Fossomatic 7 did not succeed | 13
UA Box - USB disconnect | 12
PC-006-002 | 12
Connection to UA-Interface has been lost and cannot be re-established. Please restart system. | 12
Grand Total | 8914